## Summary
Performance Characteristic | General Acceptance Criterion
| Parameter | Criterion | Ammonia | Comments
(i) Linearity | R2 – value | R2 > 0.9995 – trace analysis
|  | R2 > 0.995 – non-trace analysis | 0.9995
(ii) Sensitivity | m-value | m > 10x y-response at LOQ. m-value | 0.10002 | Sensitivity lower than expected.
|  | y-response at LOQ | 0.055796
|  | S/N Ratio > 10 | 1.3222 | S/N lower than 10
(iii) Working Range | Range | LOQ – Highest Standard | 0,3 - 10 ppm
(iv) Limit of Quantification | Value | ≤ 50 % of lowest regulatory limit defined in intended purpose | 0.3 | Met 50 % lower than SANS 241 limit.
| %RSD | < 10 % | 2.126%
| %Recovery | 90 – 110 % | 89.64%
| CRM Range | Test Result to fall in the acceptable concentration range as per CoA
(v) Limit of Detection | Value | Lowest value higher than the uncertainty of measurement. | 0.13219 | ppm
(vi) Accuracy
· Proficiency Testing | Z-scores | (-)2 to (+)2 | 0.28842 | More recent data required.
| %Compliance | >90 % | 100%
· Certified Reference Material | Range | Certified Acceptable Range
(vii) Bias
· PTS Data | Assessment of distribution | Absence of systematic bias according to QC evaluation rules | 86.67% | Z-scores above zero, possible positive bias
|  | %Error < 2 % | 2.901% | More current data required (2020)
· CRM Data | %Variance | <2 % (Reference Value against Method Value)
(viii) Precision
· Repeatability | %RSD | < 5 % (@ LOQ) | 1.722% | Good repeatability at LOQ concentration.
· Intermediate Precision | %RSD | < 10 % | 4.012% | More data required
· Reproducibility | %Compliance | > 90 % | 100% | Good reproducibility, though baised
· Regression Precision | %RSD | < 5 % | 10.12% | Acceptable variation caused by regression line at LOQ
· Injection Precision | %RSD | < 5 % | 1.515% | Good precision for responses at LOQ. (Free from calibration errors)
(ix) Robustness & Ruggedness | %RSD | Pooled %RSD ≤ 15 %
| Bias | Absence of systematic bias according to QC evaluation rules
(x) Recoveries | Percent Range | 80 – 120 % for non-trace analysis
90 – 110 % for trace analysis
(xi) Specificity and/or Selectivity | %Variance | < 5 % variance for R2 and m-values for matrix-matched and pure calibration curves
%Recovery: 80 - 120
Note on F27: USER: True test will be if the response at LOQ and the standard error are from the same curve

## Lin. & Sens
[10 empty] | STD1 |  |  |  |  | Day1
| Day 1 | Age of STDs: |  |  | Analyst: | Sthu | SC-30/01/2020 |  | R2-Values | S/N Ratio | m-Value | LOQ | Target LOQ | LOD
[8 empty] | Day 1 | 0.9999 |  | 0.0115 | 0.38532 | 0.75 | 0.1156 |  | SUMMARY OUTPUT
Calibration Results | [7 empty] | Day 2 | 0.9999 |  | 0.1482 | 0.46509 |  | 0.13953
[8 empty] | Day 3 | 0.9999 |  | 0.1508 | 0.46192 |  | 0.13857 |  | Regression Statistics
| Std. Conc. (ppm) | x | y |  |  |  |  | Day 4 | 0.9999 |  | 0.1516 | 0.45027 |  | 0.13508 |  | Multiple R | 0.99996
# |  | Calc. Conc. | Intensity | % std error |  |  |  | Day 5 | 0.9979 |  | 0.038 | 1.8716 |  | 0.56147 |  | R Square | 0.99991
1 | 0 | -0.033 | 0.008 |  |  |  |  | Day 6 | [7 empty] | Adjusted R Square | 0.9999
2 | 0.5 | 0.491 | 0.014 |  |  |  |  | Day 7 | [7 empty] | Standard Error | 0.00044447
3 | 1 | 1.115 | 0.022 |  |  |  |  | Day 8 | [7 empty] | Observations | 8
4 | 2 | 1.929 | 0.031 |  |  |  |  | Day 9
5 | 4.167 | 4.151 | 0.057 |  |  |  |  | Day 10 | [7 empty] | ANOVA
6 | 6.25 | 6.221 | 0.081 | [13 empty] | df | SS | MS | F | Significance F
7 | 8.333 | 8.421 | 0.106 |  |  |  |  | Average | 0.9995 | #DIV/0! | 0.10002 | 0.44065 |  | 0.13219 |  | Regression | 1 | 0.013458 | 0.013458 | 68125.5072 | 2.1344e-13
8 | 10 | 9.955 | 0.123 | [12 empty] | Residual | 6 | 1.1853e-06 | 1.9755e-07
[16 empty] | Total | 7 | 0.01346
[9 empty] | Average Standard Error |  |  | 0.0056485
| Day 2 | Age of STDs: |  |  | Analyst: | Kholiswa | SM-28/02/2020 | [9 empty] | Coefficients | Standard Error | t Stat | P-value | Lower 95% | Upper 95% | Lower 95,0% | Upper 95,0%
[16 empty] | Intercept | 0.0087491 | 0.00023756 | 36.8289 | 2.6738e-08 | 0.0081678 | 0.0093303 | 0.0081678 | 0.0093303
Calibration Results | [15 empty] | X Variable 1 | 0.011535 | 4.4194e-05 | 261.0086 | 2.1344e-13 | 0.011427 | 0.011643 | 0.011427 | 0.011643
|  | x | y
# | Std. Conc. (ppb) | Calc. Conc. | Intensity | % std error
1 | 0 | 0.036 | 0.022 | [11 empty] | Day2 | SUMMARY OUTPUT
2 | 0.5 | 0.533 | 0.098
3 | 1 | 0.924 | 0.159 | [12 empty] | Regression Statistics
4 | 2 | 1.982 | 0.32 | [12 empty] | Multiple R | 0.99994
5 | 4.167 | 4.198 | 0.654 | [12 empty] | R Square | 0.99987
6 | 6.25 | 6.211 | 0.951 | [12 empty] | Adjusted R Square | 0.99985
7 | 8.333 | 8.406 | 1.269 | [12 empty] | Standard Error | 0.0068929
8 | 10 | 9.96 | 1.49 | [12 empty] | Observations | 8
ANOVA
[17 empty] | df | SS | MS | F | Significance F
[16 empty] | Regression | 1 | 2.2211 | 2.2211 | 46748.7623 | 6.6046e-13
[16 empty] | Residual | 6 | 0.00028507 | 4.7512e-05
[16 empty] | Total | 7 | 2.2214
[17 empty] | Coefficients | Standard Error | t Stat | P-value | Lower 95% | Upper 95% | Lower 95,0% | Upper 95,0%
[16 empty] | Intercept | 0.02292 | 0.0036844 | 6.2208 | 0.00079725 | 0.013905 | 0.031935 | 0.013905 | 0.031935
[16 empty] | X Variable 1 | 0.14821 | 0.00068546 | 216.2146 | 6.6046e-13 | 0.14653 | 0.14988 | 0.14653 | 0.14988
| Day 3 | Age of STDs: |  |  | Analyst: | Kholiswa | 07/02/2020 | [7 empty] | Day3 | SUMMARY OUTPUT
Calibration Results | [15 empty] | Regression Statistics
[16 empty] | Multiple R | 0.99994
|  | x | y | [12 empty] | R Square | 0.99987
# | Std. Conc. (ppb) | Calc. Conc. | Intensity | % std error | [11 empty] | Adjusted R Square | 0.99985
1 | 0 | -0.017 | 0.088 | [12 empty] | Standard Error | 0.0069641
2 | 0.5 | 0.527 | 0.174 | [12 empty] | Observations | 8
3 | 1 | 1.009 | 0.249
4 | 2 | 2 | 0.403 | [12 empty] | ANOVA
5 | 4.167 | 4.116 | 0.727 | [13 empty] | df | SS | MS | F | Significance F
6 | 6.25 | 6.253 | 1.048 | [12 empty] | Regression | 1 | 2.2985 | 2.2985 | 47392.1139 | 6.3393e-13
7 | 8.333 | 8.404 | 1.365 | [12 empty] | Residual | 6 | 0.00029099 | 4.8499e-05
8 | 10 | 9.958 | 1.59 | [12 empty] | Total | 7 | 2.2987
[17 empty] | Coefficients | Standard Error | t Stat | P-value | Lower 95% | Upper 95% | Lower 95,0% | Upper 95,0%
[16 empty] | Intercept | 0.097726 | 0.0037225 | 26.2532 | 2.0152e-07 | 0.088618 | 0.10683 | 0.088618 | 0.10683
[16 empty] | X Variable 1 | 0.15077 | 0.00069255 | 217.6973 | 6.3393e-13 | 0.14907 | 0.15246 | 0.14907 | 0.15246
[15 empty] | Day4 | SUMMARY OUTPUT
| Day 4 | Age of STDs: |  |  | Analyst: | Sthu | SC-19/02/2020
Regression Statistics
Calibration Results | [15 empty] | Multiple R | 0.99994
[16 empty] | R Square | 0.99988
|  | x | y | [12 empty] | Adjusted R Square | 0.99986
# | Std. Conc. (ppb) | Calc. Conc. | Intensity | % std error | [11 empty] | Standard Error | 0.0068276
1 | 0 | -0.024 | 0.094 | [12 empty] | Observations | 8
2 | 0.5 | 0.507 | 0.178
3 | 1 | 1.017 | 0.2588 | [12 empty] | ANOVA
4 | 2 | 1.984 | 0.409 | [13 empty] | df | SS | MS | F | Significance F
5 | 4.167 | 4.264 | 0.76 | [12 empty] | Regression | 1 | 2.326 | 2.326 | 49896.3788 | 5.432e-13
6 | 6.25 | 6.095 | 1.036 | [12 empty] | Residual | 6 | 0.0002797 | 4.6617e-05
7 | 8.333 | 8.428 | 1.383 | [12 empty] | Total | 7 | 2.3263
8 | 10 | 9.979 | 1.609
[17 empty] | Coefficients | Standard Error | t Stat | P-value | Lower 95% | Upper 95% | Lower 95,0% | Upper 95,0%
[16 empty] | Intercept | 0.1047 | 0.0036491 | 28.693 | 1.1868e-07 | 0.095774 | 0.11363 | 0.095774 | 0.11363
[16 empty] | X Variable 1 | 0.15163 | 0.00067883 | 223.375 | 5.432e-13 | 0.14997 | 0.15329 | 0.14997 | 0.15329
| Day 5 | Age of STDs: |  |  | Analyst: | Sthu | SM-18/12/2019 | [7 empty] | Day 5 | SUMMARY OUTPUT
Calibration Results | [15 empty] | Regression Statistics
[16 empty] | Multiple R | 0.99896
| Std. Conc. (ppb) | x | y | [12 empty] | R Square | 0.99793
# |  | Calc. Conc. | Intensity | % std error | [11 empty] | Adjusted R Square | 0.99758
1 | 0 | -0.042 | -0.001 | [12 empty] | Standard Error | 0.0071132
2 | 0.5 | 0.525 | 0.024 | [12 empty] | Observations | 8
3 | 1 | 1.025 | 0.046
4 | 2 | 2.02 | 0.088 | [12 empty] | ANOVA
5 | 4.167 | 4.142 | 0.174 | [13 empty] | df | SS | MS | F | Significance F
6 | 6.25 | 6.191 | 0.251 | [12 empty] | Regression | 1 | 0.14608 | 0.14608 | 2887.0722 | 2.7897e-09
7 | 8.333 | 8.431 | 0.328 | [12 empty] | Residual | 6 | 0.00030358 | 5.0597e-05
8 | 10 | 9.958 | 0.377 | [12 empty] | Total | 7 | 0.14638
[17 empty] | Coefficients | Standard Error | t Stat | P-value | Lower 95% | Upper 95% | Lower 95,0% | Upper 95,0%
[16 empty] | Intercept | 0.0076627 | 0.003802 | 2.0154 | 0.090473 | -0.0016405 | 0.016966 | -0.0016405 | 0.016966
[16 empty] | X Variable 1 | 0.038006 | 0.00070734 | 53.7315 | 2.7897e-09 | 0.036275 | 0.039737 | 0.036275 | 0.039737
Note on S8: USER: outlier - rejected
Note on U8: USER: Outlier = rejected

## LOQ
Actions
•Prepare the standard (atcomputed LoQ) and analyse.
•Assess precision and accuracy at LoQ
•Recalibrate including LOQ standard, compute R2 (linearity including LoQ)
Target LOQ: | 0.75 |  |  | ppm
TRIAL 1
Rep. | Target Value
| 0.75 |  | 0.3
| ppm | Response | ppm | Response
1 | 0.75309 | 0.12839 | 0.27635 | 0.05691
2 | 0.7423 | 0.12677 | 0.26709 | 0.05552
3 | 0.77091 | 0.13106 | 0.27006 | 0.05597
4 | 0.75154 | 0.12815 | 0.27684 | 0.05698
5 | 0.76334 | 0.12992 | 0.27618 | 0.05688
6 | 0.7817 | 0.13268 | 0.26435 | 0.05511
7 | 0.75169 | 0.12818 | 0.26745 | 0.05557
8 | 0.74348 | 0.12695 | 0.26073 | 0.05457
9 | 0.759 | 0.12927 | 0.26536 | 0.05526
10 | 0.74231 | 0.12677 | 0.2649 | 0.05519
SD. | 0.013015 | 0.0019515 | 0.0057178 | 0.00085658
AV. | 0.75594 | 0.12881 | 0.26893 | 0.055796
%RSD | 1.722% | 1.515% | 2.126% | 1.535%
%Recovery | 100.8% |  | 89.64%
Date Analyesed:
Analyst
Y-Value @ LOQ (Signal) |  | 0.055796
Y-Value @ Blank (Noise) |  | 0.0422
| S/N | 1.3222
TRIAL 2
Rep. | Target Value
0.4
| ppm | Response | ppm | Response
1
2
3
4
5
6
7
8
9
10
SD. | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
AV. | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
%RSD | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0!
%Recovery | #DIV/0! |  | #DIV/0!
Date Analyesed:
Analyst

## Acc. Repro & Bais
Date | Median | Robust RSD | Lab Results | Analyst | Z-scores | %Error |  | LCL | LWL | TV | UWL | UCL
May 2018 S1 | 1.28 | 0.18 | 1.3 | David | 0.11111 | 1.563% | 1 | -3 | -2 | 0 | 2 | 3
May 2018 S2 | 25.54 | 2.52 | 27 | David | 0.57937 | 5.717% | 2 | -3 | -2 | 0 | 2 | 3
May 2018 S3 | 18.09 | 1.7 | 18.8 | David | 0.41765 | 3.925% | 3 | -3 | -2 | 0 | 2 | 3
Aug 2018 S1 | 5.44 | 0.69 | 5.64 | Kholiswa | 0.28986 | 3.676% | 4 | -3 | -2 | 0 | 2 | 3
Aug 2018 S2 | 13.97 | 1.29 | 14.05 | Kholiswa | 0.062016 | 0.5727% | 5 | -3 | -2 | 0 | 2 | 3
Aug 2018 S3 | 28.2 | 2.08 | 28.81 | Kholiswa | 0.29327 | 2.163% | 6 | -3 | -2 | 0 | 2 | 3
Nov 2018 S1 | 4.77 | 0.5 | 4.75 | David | -0.04 | -0.4193% | 7 | -3 | -2 | 0 | 2 | 3
Nov 2018 S2 | 18.4 | 1.63 | 18.5 | David | 0.06135 | 0.5435% | 8 | -3 | -2 | 0 | 2 | 3
Nov 2018 S3 | 8.99 | 0.98 | 9.38 | David | 0.39796 | 4.338% | 9 | -3 | -2 | 0 | 2 | 3
Feb 2019 S1 | 0.68 | 0.12 | 0.77 | Lungi | 0.75 | 13.24% | 10 | -3 | -2 | 0 | 2 | 3
Feb 2019 S2 | 14.08 | 1.91 | 14.59 | Lungi | 0.26702 | 3.622% | 11 | -3 | -2 | 0 | 2 | 3
Feb 2019 S3 | 19.06 | 1.96 | 19.76 | Lungi | 0.35714 | 3.673% | 12 | -3 | -2 | 0 | 2 | 3
May 2019 S1 | 0.41 | 0.09 | 0.38 | Sthu | -0.33333 | -7.317% | 13 | -3 | -2 | 0 | 2 | 3
May 2019 S2 | 20.99 | 1.57 | 21.86 | Sthu | 0.55414 | 4.145% | 14 | -3 | -2 | 0 | 2 | 3
May 2019 S3 | 9.3 | 0.68 | 9.68 | Sthu | 0.55882 | 4.086% | 15 | -3 | -2 | 0 | 2 | 3
[7 empty] | 16 | -3 | -2 | 0 | 2 | 3
|  |  |  | Average % Error |  | 2.901% | 17 | -3 | -2 | 0 | 2 | 3
|  |  |  | Average % Error |  | 0.28842
[7 empty] | 18 | -3 | -2 | 0 | 2 | 3
|  |  |  | Z-score Above 0 | 13 | 86.67% | 19 | -3 | -2 | 0 | 2 | 3
|  |  |  | Z-score Below 0 | 2 | 13.33% | 20 | -3 | -2 | 0 | 2 | 3
|  |  |  | Total | 15 |  | 21 | -3 | -2 | 0 | 2 | 3
[7 empty] | 22 | -3 | -2 | 0 | 2 | 3
|  |  |  | Outliers | 0 | 0 | 23 | -3 | -2 | 0 | 2 | 3
|  |  |  | Compliant | 15 | 100% | 24 | -3 | -2 | 0 | 2 | 3
|  |  |  | Total | 15 |  | 25 | -3 | -2 | 0 | 2 | 3
[7 empty] | 26 | -3 | -2 | 0 | 2 | 3
[7 empty] | 27 | -3 | -2 | 0 | 2 | 3
[7 empty] | 28 | -3 | -2 | 0 | 2 | 3
[7 empty] | 29 | -3 | -2 | 0 | 2 | 3
[7 empty] | 30 | -3 | -2 | 0 | 2 | 3
[7 empty] | 31 | -3 | -2 | 0 | 2 | 3
[7 empty] | 32 | -3 | -2 | 0 | 2 | 3
[7 empty] | 33 | -3 | -2 | 0 | 2 | 3
[7 empty] | 34 | -3 | -2 | 0 | 2 | 3
[7 empty] | 35 | -3 | -2 | 0 | 2 | 3
[7 empty] | 36 | -3 | -2 | 0 | 2 | 3
[7 empty] | 37 | -3 | -2 | 0 | 2 | 3
[7 empty] | 38 | -3 | -2 | 0 | 2 | 3
[7 empty] | 39 | -3 | -2 | 0 | 2 | 3
[7 empty] | 40 | -3 | -2 | 0 | 2 | 3

## Int. Prec.
Hypothetical
LIMITS
# | Date | Results | Analyst |  | LCL | LWL | TV | UWL | UCL | UCL | 5.5
1 | 27/06/2019 | 5.25 | Lungi | 1 | 4.5 | 4.65 | 5 | 5.35 | 5.5 | UWL | 5.35
2 | 03/09/2019 | 5.03 | Lungi | 2 | 4.5 | 4.65 | 5 | 5.35 | 5.5
3 | 24/06/2019 | 5.13 | Lungi | 3 | 4.5 | 4.65 | 5 | 5.35 | 5.5 | TV | 5
4 | 03/09/2019 | 5 | Lungi | 4 | 4.5 | 4.65 | 5 | 5.35 | 5.5
5 | 04/09/2019 | 5.02 | David | 5 | 4.5 | 4.65 | 5 | 5.35 | 5.5 | LWL | 4.65
6 | 11/09/2019 | 5 | Kholi | 6 | 4.5 | 4.65 | 5 | 5.35 | 5.5 | LCL | 4.5
7 | 17/09/2019 | 5.17 | Sthu | 7 | 4.5 | 4.65 | 5 | 5.35 | 5.5
8 | 17/09/2019 | 5 | Kholi | 8 | 4.5 | 4.65 | 5 | 5.35 | 5.5
9 | 18/09/2019 | 5 | Kholi | 9 | 4.5 | 4.65 | 5 | 5.35 | 5.5
10 | 25/09/2019 | 4.91 | Kholi | 10 | 4.5 | 4.65 | 5 | 5.35 | 5.5 | Calculated
11 | 27/09/2019 | 4.91 | Kholi | 11 | 4.5 | 4.65 | 5 | 5.35 | 5.5 | LIMITS
12 | 01/10/2019 | 4.81 | Sthu | 12 | 4.5 | 4.65 | 5 | 5.35 | 5.5 | UCL | 5.6972
13 | 01/10/2019 | 5.12 | Sthu | 13 | 4.5 | 4.65 | 5 | 5.35 | 5.5 | UWL | 5.4932
14 | 03/10/2019 | 4.83 | Lungi | 14 | 4.5 | 4.65 | 5 | 5.35 | 5.5
15 | 04/10/2019 | 4.73 | Sthu | 15 | 4.5 | 4.65 | 5 | 5.35 | 5.5 | TV | 5.0851
16 | 09/10/2019 | 4.73 | Lungi | 16 | 4.5 | 4.65 | 5 | 5.35 | 5.5
17 | 10/10/2019 | 5.29 | Lungi | 17 | 4.5 | 4.65 | 5 | 5.35 | 5.5 | LWL | 4.6771
18 | 14/10/2019 | 4.88 | Lungi | 18 | 4.5 | 4.65 | 5 | 5.35 | 5.5 | LCL | 4.473
19 | 15/10/2019 | 4.94 | Sthu | 19 | 4.5 | 4.65 | 5 | 5.35 | 5.5
20 | 16/10/2019 | 4.84 | Lungi | 20 | 4.5 | 4.65 | 5 | 5.35 | 5.5
21 | 16/10/2019 | 4.79 | Lungi | 21 | 4.5 | 4.65 | 5 | 5.35 | 5.5 | SD | 0.20403
22 | 23/10/2019 | 5 | Sthu | 22 | 4.5 | 4.65 | 5 | 5.35 | 5.5
23 | 23/10/2019 | 4.84 | Sthu | 23 | 4.5 | 4.65 | 5 | 5.35 | 5.5
24 | 28/10/2019 | 4.92 | Lungi | 24 | 4.5 | 4.65 | 5 | 5.35 | 5.5
25 | 05/11/2019 | 4.91 | Sthu | 25 | 4.5 | 4.65 | 5 | 5.35 | 5.5
26 | 06/11/2019 | 5.43 | Sthu | 26 | 4.5 | 4.65 | 5 | 5.35 | 5.5
27 | 06/11/2019 | 5.27 | Sthu | 27 | 4.5 | 4.65 | 5 | 5.35 | 5.5
28 | 12/11/2019 | 5.2 | Kholi | 28 | 4.5 | 4.65 | 5 | 5.35 | 5.5
29 | 13/11/2019 | 5.27 | Sthu | 29 | 4.5 | 4.65 | 5 | 5.35 | 5.5
30 | 21/11/2019 | 5.23 | Sthu | 30 | 4.5 | 4.65 | 5 | 5.35 | 5.5
31 | 21/11/2019 | 5.45 | Sthu | 31 | 4.5 | 4.65 | 5 | 5.35 | 5.5
32 | 21/11/2019 | 5.18 | Sthu | 32 | 4.5 | 4.65 | 5 | 5.35 | 5.5
33 | 21/11/2019 | 5.11 | Sthu | 33 | 4.5 | 4.65 | 5 | 5.35 | 5.5
34 | 28/11/2019 | 5.01 | Sthu | 34 | 4.5 | 4.65 | 5 | 5.35 | 5.5
35 | 28/11/2019 | 5.11 | Sthu | 35 | 4.5 | 4.65 | 5 | 5.35 | 5.5
36 | 02/12/2019 | 5.07 | Kholi | 36 | 4.5 | 4.65 | 5 | 5.35 | 5.5
37 | 03/12/2019 | 5.05 | Kholi | 37 | 4.5 | 4.65 | 5 | 5.35 | 5.5
38 | 04/12/2019 | 5.11 | Sthu | 38 | 4.5 | 4.65 | 5 | 5.35 | 5.5
39 | 04/12/2019 | 5.2 | Sthu | 39 | 4.5 | 4.65 | 5 | 5.35 | 5.5
40 | 04/12/2019 | 5.15 | Sthu | 40 | 4.5 | 4.65 | 5 | 5.35 | 5.5
41 | 04/12/2019 | 5.11 | Sthu | 41 | 4.5 | 4.65 | 5 | 5.35 | 5.5
42 | 09/12/2019 | 5.01 | Kholi | 42 | 4.5 | 4.65 | 5 | 5.35 | 5.5
43 | 10/12/2019 | 5.01 | Kholi | 43 | 4.5 | 4.65 | 5 | 5.35 | 5.5
44 | 11/12/2019 | 5 | Kholi | 44 | 4.5 | 4.65 | 5 | 5.35 | 5.5
45 | 12/12/2019 | 5.08 | Sthu | 45 | 4.5 | 4.65 | 5 | 5.35 | 5.5
46 | 12/12/2019 | 5 | Sthu | 46 | 4.5 | 4.65 | 5 | 5.35 | 5.5
47 | 18/12/2019 | 5.25 | Sthu | 47 | 4.5 | 4.65 | 5 | 5.35 | 5.5
48 | 07/01/2020 | 5.22 | Kholi | 48 | 4.5 | 4.65 | 5 | 5.35 | 5.5
49 | 08/01/2020 | 5.12 | Kholi | 49 | 4.5 | 4.65 | 5 | 5.35 | 5.5
50 | 15/01/2020 | 5.02 | Kholi | 50 | 4.5 | 4.65 | 5 | 5.35 | 5.5
51 | 21/01/2019 | 5.36 | Sthu | 51 | 4.5 | 4.65 | 5 | 5.35 | 5.5
52 | 21/01/2020 | 5.39 | Sthu | 52 | 4.5 | 4.65 | 5 | 5.35 | 5.5
53 | 27/01/2020 | 5.02 | Kholi | 53 | 4.5 | 4.65 | 5 | 5.35 | 5.5
54 | 28/01/2020 | 4.81 | Sthu | 54 | 4.5 | 4.65 | 5 | 5.35 | 5.5
55 | 28/01/2020 | 5.02 | Sthu | 55 | 4.5 | 4.65 | 5 | 5.35 | 5.5
56 | 29/01/2020 | 5.34 | Sthu | 56 | 4.5 | 4.65 | 5 | 5.35 | 5.5
57 | 04/02/2020 | 4.58 | Sthu | 57 | 4.5 | 4.65 | 5 | 5.35 | 5.5
58 | 04/02/2020 | 4.64 | Sthu | 58 | 4.5 | 4.65 | 5 | 5.35 | 5.5
59 | 04/02/2020 | 4.57 | Sthu | 59 | 4.5 | 4.65 | 5 | 5.35 | 5.5
60 | 07/02/2020 | 5.02 | Kholi | 60 | 4.5 | 4.65 | 5 | 5.35 | 5.5
61 | 11/02/2020 | 5.36 | Sthu | 61 | 4.5 | 4.65 | 5 | 5.35 | 5.5
62 | 11/02/2020 | 5.3 | Sthu | 62 | 4.5 | 4.65 | 5 | 5.35 | 5.5
63 | 19/02/2020 | 5.3 | Kholi | 63 | 4.5 | 4.65 | 5 | 5.35 | 5.5
64 | 19/02/2020 | 5.1 | Kholi | 64 | 4.5 | 4.65 | 5 | 5.35 | 5.5
65 | 20/02/2020 | 5.47 | Sthu | 65 | 4.5 | 4.65 | 5 | 5.35 | 5.5
66 | 20/02/2020 | 5.28 | Sthu | 66 | 4.5 | 4.65 | 5 | 5.35 | 5.5
67 | 20/02/2020 | 5.26 | Sthu | 67 | 4.5 | 4.65 | 5 | 5.35 | 5.5
68 | 26/02/2020 | 5.12 | Kholi | 68 | 4.5 | 4.65 | 5 | 5.35 | 5.5
69 | 26/02/2020 | 5.15 | Kholi | 69 | 4.5 | 4.65 | 5 | 5.35 | 5.5
70 | 04/03/2020 | 5.26 | Lungi | 70 | 4.5 | 4.65 | 5 | 5.35 | 5.5
71 | 11/03/2020 | 5.35 | Lungi | 71 | 4.5 | 4.65 | 5 | 5.35 | 5.5
72 | 16/03/2020 | 5.29 | Lungi | 72 | 4.5 | 4.65 | 5 | 5.35 | 5.5
73 | 18/03/2020 | 5.3 | Lungi | 73 | 4.5 | 4.65 | 5 | 5.35 | 5.5
74 | 19/03/2020 | 5.3 | Lungi | 74 | 4.5 | 4.65 | 5 | 5.35 | 5.5
75 |  |  |  | 75 | 4.5 | 4.65 | 5 | 5.35 | 5.5
76 |  |  |  | 76 | 4.5 | 4.65 | 5 | 5.35 | 5.5
77 |  |  |  | 77 | 4.5 | 4.65 | 5 | 5.35 | 5.5
78 |  |  |  | 78 | 4.5 | 4.65 | 5 | 5.35 | 5.5
79 |  |  |  | 79 | 4.5 | 4.65 | 5 | 5.35 | 5.5
80 |  |  |  | 80 | 4.5 | 4.65 | 5 | 5.35 | 5.5
81 |  |  |  | 81 | 4.5 | 4.65 | 5 | 5.35 | 5.5
82 |  |  |  | 82 | 4.5 | 4.65 | 5 | 5.35 | 5.5
83 |  |  |  | 83 | 4.5 | 4.65 | 5 | 5.35 | 5.5
84 |  |  |  | 84 | 4.5 | 4.65 | 5 | 5.35 | 5.5
85 |  |  |  | 85 | 4.5 | 4.65 | 5 | 5.35 | 5.5
86 |  |  |  | 86 | 4.5 | 4.65 | 5 | 5.35 | 5.5
87 |  |  |  | 87 | 4.5 | 4.65 | 5 | 5.35 | 5.5
88 |  |  |  | 88 | 4.5 | 4.65 | 5 | 5.35 | 5.5
89 |  |  |  | 89 | 4.5 | 4.65 | 5 | 5.35 | 5.5
90 |  |  |  | 90 | 4.5 | 4.65 | 5 | 5.35 | 5.5
91 |  |  |  | 91 | 4.5 | 4.65 | 5 | 5.35 | 5.5
92 |  |  |  | 92 | 4.5 | 4.65 | 5 | 5.35 | 5.5
93 |  |  |  | 93 | 4.5 | 4.65 | 5 | 5.35 | 5.5
94 |  |  |  | 94 | 4.5 | 4.65 | 5 | 5.35 | 5.5
95 |  |  |  | 95 | 4.5 | 4.65 | 5 | 5.35 | 5.5
96 |  |  |  | 96 | 4.5 | 4.65 | 5 | 5.35 | 5.5
97 |  |  |  | 97 | 4.5 | 4.65 | 5 | 5.35 | 5.5
98 |  |  |  | 98 | 4.5 | 4.65 | 5 | 5.35 | 5.5
99 |  |  |  | 99 | 4.5 | 4.65 | 5 | 5.35 | 5.5
100 |  |  |  | 100 | 4.5 | 4.65 | 5 | 5.35 | 5.5
101 |  |  |  | 101 | 4.5 | 4.65 | 5 | 5.35 | 5.5
102 |  |  |  | 102 | 4.5 | 4.65 | 5 | 5.35 | 5.5
103 |  |  |  | 103 | 4.5 | 4.65 | 5 | 5.35 | 5.5
104 |  |  |  | 104 | 4.5 | 4.65 | 5 | 5.35 | 5.5
105 |  |  |  | 105 | 4.5 | 4.65 | 5 | 5.35 | 5.5
106 |  |  |  | 106 | 4.5 | 4.65 | 5 | 5.35 | 5.5
107 |  |  |  | 107 | 4.5 | 4.65 | 5 | 5.35 | 5.5
108 |  |  |  | 108 | 4.5 | 4.65 | 5 | 5.35 | 5.5
109 |  |  |  | 109 | 4.5 | 4.65 | 5 | 5.35 | 5.5
110 |  |  |  | 110 | 4.5 | 4.65 | 5 | 5.35 | 5.5
111 |  |  |  | 111 | 4.5 | 4.65 | 5 | 5.35 | 5.5
112 |  |  |  | 112 | 4.5 | 4.65 | 5 | 5.35 | 5.5
113 |  |  |  | 113 | 4.5 | 4.65 | 5 | 5.35 | 5.5
114 |  |  |  | 114 | 4.5 | 4.65 | 5 | 5.35 | 5.5
115 |  |  |  | 115 | 4.5 | 4.65 | 5 | 5.35 | 5.5
116 |  |  |  | 116 | 4.5 | 4.65 | 5 | 5.35 | 5.5
117 |  |  |  | 117 | 4.5 | 4.65 | 5 | 5.35 | 5.5
118 |  |  |  | 118 | 4.5 | 4.65 | 5 | 5.35 | 5.5
119 |  |  |  | 119 | 4.5 | 4.65 | 5 | 5.35 | 5.5
120 |  |  |  | 120 | 4.5 | 4.65 | 5 | 5.35 | 5.5
121 |  |  |  | 121 | 4.5 | 4.65 | 5 | 5.35 | 5.5
122 |  |  |  | 122 | 4.5 | 4.65 | 5 | 5.35 | 5.5
123 |  |  |  | 123 | 4.5 | 4.65 | 5 | 5.35 | 5.5
124 |  |  |  | 124 | 4.5 | 4.65 | 5 | 5.35 | 5.5
125 |  |  |  | 125 | 4.5 | 4.65 | 5 | 5.35 | 5.5
126 |  |  |  | 126 | 4.5 | 4.65 | 5 | 5.35 | 5.5
127 |  |  |  | 127 | 4.5 | 4.65 | 5 | 5.35 | 5.5
128 |  |  |  | 128 | 4.5 | 4.65 | 5 | 5.35 | 5.5
129 |  |  |  | 129 | 4.5 | 4.65 | 5 | 5.35 | 5.5
130 |  |  |  | 130 | 4.5 | 4.65 | 5 | 5.35 | 5.5
131 |  |  |  | 131 | 4.5 | 4.65 | 5 | 5.35 | 5.5
132 |  |  |  | 132 | 4.5 | 4.65 | 5 | 5.35 | 5.5
133 |  |  |  | 133 | 4.5 | 4.65 | 5 | 5.35 | 5.5
134 |  |  |  | 134 | 4.5 | 4.65 | 5 | 5.35 | 5.5
135 |  |  |  | 135 | 4.5 | 4.65 | 5 | 5.35 | 5.5
136 |  |  |  | 136 | 4.5 | 4.65 | 5 | 5.35 | 5.5
137 |  |  |  | 137 | 4.5 | 4.65 | 5 | 5.35 | 5.5
138 |  |  |  | 138 | 4.5 | 4.65 | 5 | 5.35 | 5.5
139 |  |  |  | 139 | 4.5 | 4.65 | 5 | 5.35 | 5.5
140 |  |  |  | 140 | 4.5 | 4.65 | 5 | 5.35 | 5.5
141 |  |  |  | 141 | 4.5 | 4.65 | 5 | 5.35 | 5.5
142 |  |  |  | 142 | 4.5 | 4.65 | 5 | 5.35 | 5.5
143 |  |  |  | 143 | 4.5 | 4.65 | 5 | 5.35 | 5.5
144 |  |  |  | 144 | 4.5 | 4.65 | 5 | 5.35 | 5.5
145 |  |  |  | 145 | 4.5 | 4.65 | 5 | 5.35 | 5.5
146 |  |  |  | 146 | 4.5 | 4.65 | 5 | 5.35 | 5.5
147 |  |  |  | 147 | 4.5 | 4.65 | 5 | 5.35 | 5.5
148 |  |  |  | 148 | 4.5 | 4.65 | 5 | 5.35 | 5.5
149 |  |  |  | 149 | 4.5 | 4.65 | 5 | 5.35 | 5.5
150 |  |  |  | 150 | 4.5 | 4.65 | 5 | 5.35 | 5.5
151 |  |  |  | 151 | 4.5 | 4.65 | 5 | 5.35 | 5.5
152 |  |  |  | 152 | 4.5 | 4.65 | 5 | 5.35 | 5.5
153 |  |  |  | 153 | 4.5 | 4.65 | 5 | 5.35 | 5.5
154 |  |  |  | 154 | 4.5 | 4.65 | 5 | 5.35 | 5.5
155 |  |  |  | 155 | 4.5 | 4.65 | 5 | 5.35 | 5.5
156 |  |  |  | 156 | 4.5 | 4.65 | 5 | 5.35 | 5.5
157 |  |  |  | 157 | 4.5 | 4.65 | 5 | 5.35 | 5.5
158 |  |  |  | 158 | 4.5 | 4.65 | 5 | 5.35 | 5.5
159 |  |  |  | 159 | 4.5 | 4.65 | 5 | 5.35 | 5.5
160 |  |  |  | 160 | 4.5 | 4.65 | 5 | 5.35 | 5.5
|  |  |  | 161 | 4.5 | 4.65 | 5 | 5.35 | 5.5
| TV | 5
| Average | 5.0851 |  | 162 | 4.5 | 4.65 | 5 | 5.35 | 5.5
| SD | 0.20403 |  | 163 | 4.5 | 4.65 | 5 | 5.35 | 5.5
| %RSD | 4.012%
| %Error | 1.703%
| Outliers | 0 | 0
| Compliant | 15 | 100%
| Total | 15

## Recoveries
| ACTION: | Prepare calibration standards from metrologically traceable material in ultrapure water - calibrate equipment. Substitute ultrapure water with the different sample matrices
and prepare calibration standards following similar procedure - run these solutions as "sample". Plot all the results on the same cartesian plan and compute R2 and m-values
Day 1
Calibration Results for "Ultrapure Water" |  |  |  |  | Calibration Results for "Matrice 1: Drinking Water (<1 NTU)"
|  | X | Y |  |  |  | X (Concentration, ppm) |  |  | Y (Response)
# | Std. Conc. (ppb) | Calc. Conc. | Response | % Std Error | # | Spike Conc (ppm) | Rep 1 | Rep 2 | Rep 3 | Rep 1 | Rep 2 | Rep 3
1 | Blank | -0.04717 | 0.01717 |  | 1 | Unspiked | -0.04676 | -0.05725 | -0.05845 | 0.01724 | 0.01564 | 0.01546
3 | 0.5 | 0.69787 | 0.12946 |  | 2 | 0.3 | 0.26923 | 0.24115 | 0.24613 | 0.0651 | 0.06086 | 0.06161
4 | 1 | 0.83598 | 0.15006 |  | 3 | 1 | 0.98405 | 0.96047 | 0.95045 | 0.17207 | 0.16857 | 0.16708
5 | 2 | 2.1196 | 0.33817 |  | 4 | 4 | 4.1385 | 4.1352 | 4.164 | 0.62193 | 0.62147 | 0.62542
6 | 4.167 | 3.8133 | 0.57722 |  | 5 | 8 | 8.8277 | 8.9644 | 8.7774 | 1.2239 | 1.2402 | 1.2178
7 | 6.25 | 6.5538 | 0.94194 |  | 6 | 10 | 8.677 | 8.9493 | 8.7791 | 0.15478 | 0.15883 | 0.1563
8 | 8.333 | 8.3727 | 1.1689
9 | 10 | 9.9077 | 1.3512
Calibration Results for "Matrice 2: Surface Water (>5 NTU)"
[7 empty] | X (Concentration, ppm) |  |  | Y (Response)
[5 empty] | # | Spike Conc (ppm) | Rep 1 | Rep 2 | Rep 3 | Rep 1 | Rep 2 | Rep 3
[5 empty] | 1 | Unspiked | -0.04323 | -0.04726 | -0.04221 | 0.01777 | 0.01716 | 0.01793
[5 empty] | 2 | 0.3 | 0.2755 | 0.24675 | 0.24399 | 0.06605 | 0.06171 | 0.06129
[5 empty] | 3 | 1 | 0.95126 | 0.94996 | 0.9704 | 0.1672 | 0.16701 | 0.17004
[5 empty] | 4 | 4 | 4.42 | 4.5 | 4.43 | 0.644 | 0.655 | 0.645
[5 empty] | 5 | 8 | 8.8733 | 8.878 | 8.9243 | 1.2293 | 1.2299 | 1.2354
[5 empty] | 6 | 10 | 8.1636 | 9.9352 | 8.7486 | 0.14714 | 0.17347 | 0.15585
Calibration Results for "Matrice 3: Trade Effluent/Wastewater"
[7 empty] | X (Concentration, ppm) |  |  | Y (Response)
[5 empty] | # | Spike Conc (ppm) | Rep 1 | Rep 2 | Rep 3 | Rep 1 | Rep 2 | Rep 3
[5 empty] | 1 | Unspiked | 20.1932 | 20.1839 | 19.7477 | 0.32369 | 0.32356 | 0.31725
[5 empty] | 2 | 0.3 | 19.4525 | 17.9488 | 18.6092 | 0.31297 | 0.29115 | 0.30075
[5 empty] | 3 | 1 | 18.7322 | 18.6278 | 17.7435 | 0.30253 | 0.30102 | 0.28817
[5 empty] | 4 | 4 | 21.432 | 21.9717 | 21.6484 | 0.34158 | 0.34935 | 0.34469
[5 empty] | 5 | 8 | 25.6313 | 24.9564 | 25.3466 | 0.40178 | 0.39215 | 0.39772
[5 empty] | 6 | 10 | 27.1246 | 27.1324 | 26.7532 | 0.42304 | 0.42315 | 0.41776

## Matrices Cal
| ACTION: | Prepare calibration standards from metrologically traceable material in ultrapure water - calibrate equipment. Substitute ultrapure water with the different sample matrices
and prepare calibration standards following similar procedure - run these solutions as "sample". Plot all the results on the same cartesian plan and compute R2 and m-values
Day 1
Calibration Results for "Ultrapure Water" |  |  |  |  | Calibration Results for "Matrice 1: Drinking Water" |  |  |  |  | Calibration Results for "Matrice 2: Surface Water (>5 NTU)" |  |  |  |  | Calibration Results for "Matrice 3: Trade Effluent"
|  | x | y |  |  |  | x | y |  |  |  | x | y |  |  |  | x | y
# | Std. Conc. (ppb) | Calc. Conc. | Response | % std error | # | Std. Conc. | Calc. Conc. | Response | % std error | # | Std. Conc. | Calc. Conc. | Response | % std error | # | Std. Conc. | Calc. Conc. | Response | % std error
1 | 0 |  |  |  | 1 |  |  |  |  | 1 |  |  |  |  | 1
2 | 0.005 |  |  |  | 2 |  |  |  |  | 2 |  |  |  |  | 2
3 | 0.01 |  |  |  | 3 |  |  |  |  | 3 |  |  |  |  | 3
4 | 0.02 |  |  |  | 4 |  |  |  |  | 4 |  |  |  |  | 4
5 | 0.04 |  |  |  | 5 |  |  |  |  | 5 |  |  |  |  | 5
6 | 0.06 |  |  |  | 6 |  |  |  |  | 6 |  |  |  |  | 6
7 | 0.08
[5 empty] | Unspiked sample concentration | [5 empty] | Unspiked sample concentration |  |  |  |  | Unspiked sample concentration
Calibration Results for "Ultrapure Water"
|  | x | y
# | Std. Conc. (ppb) | Calc. Conc. | Response | % std error
1 | 0
2 | 0.005
3 | 0.01
4 | 0.02
5 | 0.04
6 | 0.06
7 | 0.08

## Limits
Surface Water Lake Cubhu and Mzingazi
Determinants | Units | Class I | Class II | Class III
Arsenic as As | mg/L | 0 – 0.01 | 0.01 – 0.2 | £0.2 – 0.3
Cadmium | mg/L | 0 – 0.005 | £0.005 | £ 0.005
Copper as Cu | mg/L | 0 – 1.0 | 1 – 3 | £3 – 30
Iron as Fe | mg/L | 0 – 0.1 | 0.1 – 0.3 | £0.3 – 1.0
Lead as Pb | mg/L | 0 – 0.01 | 0.01 | £ 0.01
Manganese as Mn | mg/L | 0 – 0.05 | 0.05 – 0.1 | £0.1 – 0.15
Mercury as Hg | mg/L | 0 – 0.001 | 0.001 | £ 0.001
Selenium as Se | mg/L | 0 – 0.02 | 0.02 – 0.05 | £0.05 – 0.1
Zinc as Zn | mg/L | 0 – 3 | 3 – 5 | £5 – 10
Surface Water Storm Water (Industrial) and Pump-stations
Determinants | Units | Max Allow.
Arsenic as As | mg/L | 0.1
Boron as B | mg/L | 0.5
Cadmium as Cd | mg/L | 0.05
Copper as Cu | mg/L | 0.02
Lead as Pb | mg/L | 0.1
Iron as Fe | mg/L | 0.3
Manganese as Mn | mg/L | 0.1
Mercury as Hg | mg/L | 0.02
Selenium as Se | mg/L | 0.05
Total Chromium as Cr | mg/L | 0.5
Zinc as Zn | mg/L | 0.3
Ground Water - Cemetries
Determinants | Units | Max Allow.
Arsenic as As | mg/L | 0.1
Boron as B | mg/L | 0.5
Cadmium as Cd | mg/L | 0.05
Copper as Cu | mg/L | 0.02
Lead as Pb | mg/L | 0.1
Iron as Fe | mg/L | 0.3
Manganese as Mn | mg/L | 0.1
Mercury as Hg | mg/L | 0.02
Nickel as Ni | mg/L | 50
Selenium as Se | mg/L | 0.05
Total Chromium as Cr | mg/L | 0.05
Zinc as Zn | mg/L | 0.3
Ground Water - Landfill Sites
Determinants | Units | Max Allow.
Arsenic as As | mg/L | 0.1
Boron as B | mg/L | 0.5
Cadmium as Cd | mg/L | 0.05
Copper as Cu | mg/L | 0.02
Lead as Pb | mg/L | 0.1
Iron as Fe | mg/L | 0.3
Manganese as Mn | mg/L | 0.1
Mercury as Hg | mg/L | 0.02
Selenium as Se | mg/L | 0.05
Total Chromium as Cr | mg/L | 0.05
Zinc as Zn | mg/L | 0.3
Industrial Effluent - Sch B
General Quality Limits | Large Works | Small Works | Units
| > 25 Ml/d | < 25 Ml/d
Copper (as Cu) | 50 | 5 | mg/l
Nickel (Ni) | 50 | 5 | mg/l
Zinc (Zn) | 50 | 5 | mg/l
Iron (Fe) | 50 | 5 | mg/l
Boron (B) | 50 | 5 | mg/l
Selenium (Se) | 50 | 5 | mg/l
Manganese (Mn) | 50 | 5 | mg/l
Lead (Pb) | 20 | 5 | mg/l
Cadmium (Cd) | 20 | 5 | mg/l
Mercury (Hg) | 1 | 1 | mg/l
Total Chromium (Cr) | 20 | 5 | mg/l
Arsenic (As) | 20 | 5 | mg/l
Titanium (Ti) | 20 | 5 | mg/l
Cobalt (Co) | 20 | 5 | mg/l
Total Metals | 100 | 20 | mg/l
Industrial Effluent - Sch C
Determinants | Limits | Units
Arsenic (expressed as As) | 5 | mg/l
Cadmium (expressed as Cd) | 1,5 | mg/l
Total chromium (expressed as Cr) | 3 | mg/l
Copper (expressed as Cu) | 3 | mg/l
Lead (expressed as Pb) | 5 | mg/l
Mercury (expressed as Hg) | 0,05 | mg/l
Nickel (expressed as Ni) | 10 | mg/l
Zinc (expressed as Zn) | 20 | mg/l
[5 empty] | SANS241 |  | Desired LOQ | Target LOQ
Ammonia |  |  |  |  | 1.5 | ppm | 0.5 | 0.75 | ppm